## U with Nordtest, Fe in seawater
Measurement uncertainty, Nordtest approach |  | Within-lab reproducibility data, for u(Rw) | [5 empty] | Bias data, for u(bias)
Fe in seawater |  | Five samples, measured on different days | [5 empty] | Certified reference materials D1 and D2
Data from: https://doi.org/10.3389/fmars.2018.00515
uncertainty is estimated in relative quantities |  |  | GSC (nM) | D2 (nM) | D1 (nM) | CSIS (nM) | SCIS (nM) |  | D1 | D2
|  | Mean | 1.4863 | 0.95561 | 0.69389 | 1.121 | 1.154 | Fe (nM) | 0.68675 | 0.95632
Results: |  | SD | 0.14494 | 0.10727 | 0.035389 | 0.079171 | 0.081904 | U (k = 2) (nM) | 0.08 | 0.046
C(Fe) (nM) = | 1.23 | RSD | 0.097516 | 0.11225 | 0.051001 | 0.070623 | 0.070977 | uc (nM) | 0.04 | 0.023
uc (rel) = | 0.095146 | RSD^2 | 0.0095093 | 0.0126 | 0.0026011 | 0.0049876 | 0.0050377 | uc (rel) | 0.058245 | 0.02405
uc (nM) = | 0.11703 | n-1 | 14 | 13 | 3 | 30 | 17 | Lab result | 0.69389 | 0.95561
U (k = 2) (nM) = | 0.23406 | [6 empty] | biasi (nM) | 0.0071429 | -0.0007103
|  | u(Rw) (rel) = pooled RSD = |  |  | 0.083744 |  |  | biasi (rel) | 0.010401 | -0.00074274
[8 empty] | u(Cref) (rel) = | 0.044559
|  | Date | GSC (nM) | D2 (nM) | D1 (nM) | CSIS (nM) | SCIS (nM) | RMSbias (rel) = | 0.0073734
|  | 2014-07-07 |  | 0.91261 |  |  |  | u(bias) (rel) = | 0.045165
|  | 2014-07-09 |  | 0.84427
|  | 2014-07-10 | 1.3757 | 0.87778
|  | 2014-07-11 | 1.1366 | 0.86899
|  | 2014-07-13 | 1.2588 | 1.0392
|  | 2014-07-29 | 1.4115 | 0.89357
1.4706
|  | 2015-01-20 |  | 0.90924
|  | 2015-01-23 | 1.5584
|  | 2015-01-25 | 1.5944 |  |  | 0.98303
|  | 2015-01-27 | 1.5536 |  |  | 1.0796
|  | 2015-02-05 | 1.5744 |  |  | 1.0409
|  | 2015-02-07 | 1.6235 |  |  | 1.0489
|  | 2015-02-10 |  |  |  | 1.0905
|  | 2015-02-11 |  |  |  | 1.0458
|  | 2015-02-13 |  |  |  | 1.1141
|  | 2015-02-20 |  |  |  | 1.1563
|  | 2015-02-25 |  |  |  | 1.191
|  | 2015-03-03 |  |  |  | 1.134
|  | 2015-03-04 |  |  |  | 1.1153
|  | 2015-03-06 |  |  |  | 1.345
|  | 2015-03-09 |  |  |  | 1.0677
|  | 2015-03-10 |  | 0.94607 |  | 1.0928
|  | 2015-05-07 |  | 0.8463 |  | 0.95151
|  | 2015-05-11 | 1.5897 |  |  | 1.2016
|  | 2015-05-12 | 1.4189 |  |  | 1.1067
|  | 2015-05-14 | 1.6716 |  | 0.73125 | 1.1295
|  | 2015-05-18 |  |  | 0.65108 | 1.1792
|  | 2015-05-19 |  |  | 0.68059 | 1.1544
|  | 2015-06-09 |  |  |  | 1.0803
|  | 2015-06-12 |  |  |  | 1.112
1.14
|  | 2015-09-08 | 1.5074 | 0.99782 |  | 1.1998
|  | 2015-09-10 | 1.5495 | 1.0376 |  | 1.1485
|  | 2015-09-14 |  |  |  | 1.1654
|  | 2015-09-17 |  |  |  | 1.1573 | 1.0875
|  | 2015-09-21 |  |  | 0.71264 | 1.1638 | 1.1727
|  | 2015-09-25 |  |  |  | 0.97882 | 1.0265
|  | 2015-09-26 |  |  |  | 1.24 | 1.1426
|  | 2015-09-28 |  |  |  | 1.1382 | 1.1368
|  | 2015-10-06 |  |  |  |  | 1.1108
|  | 2015-10-08 |  |  |  |  | 1.05
|  | 2015-10-13 |  |  |  |  | 1.07
|  | 2015-11-02 |  |  |  |  | 1.3114
|  | 2015-11-03 |  |  |  |  | 1.1647
|  | 2015-11-04 |  |  |  |  | 1.1976
|  | 2015-11-11 |  |  |  |  | 1.11
|  | 2016-04-06 |  | 0.96 |  |  | 1.16
|  | 2016-04-18 |  | 0.9952 |  |  | 1.2155
|  | 2016-04-19 |  | 1.25 |  |  | 1.17
|  | 2016-04-29 |  |  |  |  | 1.26
|  | 2016-06-21 |  |  |  |  | 1.3063
|  | 2016-06-23 |  |  |  |  | 1.0788